**3. Instrument Details**

**3.1. Instrument Manufacturer**

Beckman Coulter

 http://www.beckmancoulter.com/

**3.2. Instrument Model**

MoFlo Astrios integrated to Baker Hood

 Serial number: AV12013

 Manufactured: March 2013

Pre-shipment QC: completed before July 2013 (Moflo), 04/26/13 (Baker Hood)

 **3.3. Instrument Configuration and Settings**

 **3.3.1. Fluidics**

 This is a jet-in-air cell sorter.

 **3.3.2. Light Sources**

 Pinhole 1: 640-nm Newport Corporation, model EXLR-640C-105-219, solid state; 105mW

 Pinhole 2: 488-nm Coherent, solid state; 200mW

 Pinhole 3: empty

 Pinhole 4: 561-nm MPB Communications, model VFL560V3, solid state; 200mW

 Pinhole 5: empty

 Pinhole 6: 405-nm Newport Corporation, model EXLR-405C-55-219, solid state; 55mW

 Pinhole 7: 355-nm JDSU solid state; 100mW, SNF608364

 **3.3.3. Excitation Optics Configuration**

The Laser engine houses all lasers except for the 355nm. They are delivered to pinholes via fiber optic and their alignment in relation to one another is fixed. These lasers utilize top-hat optics. The UV laser is stand-alone laser, using Gaussian Beam Shaping Optics and focused manually to the pinhole position.

 **3.3.4. Optics Filter**

 The instrument received the EQ upgrade on February 24, 2015, which included dual PMTs for FSC discrimination and a 200mW laser for 488nm.

 The instrument can accept optical configuration changes. Refer to the attached configurations found in 3.3.5 Optical Detectors

 **3.3.5. Optical Detectors**

Installation configuration is listed below for each laser detection pod.

 Optical Detectors are modular.

 







**3.3.6. Optical Paths**

 Optical filters are removable.

**3.4. Other Relevant Instrument Details**

* MoFloAstrios User’s Guide: PN# B22986C Astrios IFU
* Summit Software v.6.2.7.16492 installed 2/24/15